## Slide 1: Tokozások termikus tesztje, minősítése

## Slide 2: Termikus tranziens tesztelés
The measured a(t) response function is characteristic to the package.
The features of the chip + package + environment structure can be extracted from it.

## Slide 3: Az ún. struktúrafüggvények
Measured thermal impedance curve
Derivative of the thermal impedance curve
Numerical derivation
Time-constant spectrum
Numerical deconvolution
Discretization
Foster model of the impedance
Transformation
Cauer model of the impedance
Structure function

## Slide 4: Struktúrafüggvény:
ambient
junction
ambient

## Slide 5: Struktúrafüggvény
Cold-plate
Base (Cu)
Chip (Si)
Die attach
Grease
heat-flow
Chip: small Rth/Cth ratio
Die attach: large Rth/Cth ratio
Base: small Rth/Cth ratio
Grease: large Rth/Cth ratio
Cold-plate: infinite heat-sinking (Cth  )
Junction
t
P(t)
T(z)
z = ln t
We measure the thermal transient at the junction...
...and we convert it into the cumulative structure function and a compact model:
Junction: is always in the origin
thermal Capacitance vs. thermal Resistance
Allows structural analysis and modeling…

## Slide 6: Chip rögzítés minősítése
Cold-plate
Base
Chip
Die attach
Junction
Grease
Base
Grease
Die attach
Chip
Reference device with good DA
Cold-plate
Base
Chip
Die attach
Junction
Grease
Unknown device with suspected DA voids
Copy the reference structure function into this plot
This increase suggests DA voids
Identify its structure function:
Identify its structure function:
Notes: Easy to compare values with against what they should be.
In die attach analysis a reference should be established (left hand side). Any difference with respect to the established reference indicates a failure.
The location of the failure is also identified this way.

## Slide 7: Példa: tranzisztor modul
Thermal transient testing of IC packages
Identification of properties of the chip-to-ambient heat-conduction path based on thermal transient measurements
The so called structure functions provide the required information
0
5
10
15
20
25
30
35
1e-06
1e-05
0.0001
0.001
0.01
0.1
1
10
100
Temperature rise [Cels]
Time [s]
T3Ster: RECORD=B1D
"B01d.MR1"
"B02d.mr1"
"B03d.mr1"
"B04d.mr1"
"B05d.mr1"
"B06d.mr1"
0.01
0.1
1
10
100
1000
0
1
2
3
4
5
K [W2s/K2]
Rth [K/W]
T3Ster: differential structure function
"C08D.stf"
"C02D.stf"
"C05D.stf"
"C17D.stf"

## Slide 8
The "good" structure function
Rth=3.2 K/W
Measurement of a power BJT mount: failure analysis
Notes: A függvénygörbe bal oldali vége a chipnek felel meg, a jobb oldali a cold plate-nek. Utóbbi pontot a 4 jelű nyíl mutatja. A kettő között a vízszintes tengelyen leolvasott hőellenállás a chip és a cold plate közötti teljes hőellenállás (jelen esetben 3,2 K/W).
A függvény kezdete kissé zajos (cikk-cakkos), de átlag értékként kb. K=0,1 rendelhető hozzá. Ez átszámolva szilícium keresztmetszetre 19,7 mm2 keresztmetszetet ad, ami a teljesítmény tranzisztor chip felületére reális adat.
A következő, 1 jelű csúcs a tranzisztor-tok hőkapacitása (azon belül is a tok Cu alaplemezének domináns hőkapacitása). A 2 kisebb csúcs (inkább csak dudor) a szerelőlemezen levő Cu sziget hőkapacitása, míg a 3 csúcs a szerelőlemez hőkapacitása.
Ezen pontok behatárolása után leolvashatjuk a rész-hőellenállásokat. Az 1 és 2 pont közötti kb. 0,6 K/W a tranzisztor felforrasztásának hőellenállása. A 2 és 3 pont között a szerelőlemez műanyag bevonatának hőellenállása van, jelenleg kb. 1,3 K/W. Végül a 3 és 4 pontok között a szerelőlemez és a cold plate között megvalósított hőellenállást olvashatjuk le.

## Slide 9
Measurement of a power BJT mount: failure analysis
Die attach delamination inside the package
Imperfect soldering of the package
Notes: Az 1 minimum C17-nél kb. 0,4 K/W értékkel jobbra tolódva jelenik meg, és a görbe egész további jobb oldali része nagyjából ugyanekkora eltolódást mutat. Ez egyértelműen azt mutatta, hogy a chip és a tok (1-el jelölt) Cu platformja között anomálisan nagy hőellenállás van, vagyis a chip nincs jól felforrasztva a tok belsejében, felvált a platformról! Ez az úgynevezett die attach hiba, ami megbízhatósági szempontból meglehetősen veszélyes.

## Slide 10: Két működő chip 44L LLP tokban
Top die
2.54×2.54 mm
Bottom die
3.81×3.81 mm
Mould
Leadframe
2nd die attach
1st die attach

## Slide 11
For the last 15 K/W section the thermal wave propagates in the far region of the board and in the chamber,
the structures fit perfectly (the blue curve is the shifted copy of the black curve)
Stacked die tokok tesztelése
Failure between package body and the test board: solder delamination
We can see what was invisible even for X-ray microscopy. Non-destructive analysis without using extra elements: delamination/voiding in the 2nd DA